Nano-functionalized Paper-based IoT enabled Devices for Point-of-Care Testing: A Review

Vinay Kishnani1, Sungjune Park2, Umesh T. Nakate2, Kunal Mondal3, Ankur Gupta1,*

Abstract- Over the last few years, the microfluidics phenomenon coupled with the Internet of Things (IoT) using innovative nano-functional materials have been recognized as one of the sustainable and economical tools for point of care testing (POCT) of various pathogens influencing the human health. The sensors based on these phenomena are aimed to be designed for their cost-effectiveness, to make it handy, environment-friendly and to get an accurate, easy, and rapid response. Considering the burgeoning importance of analytical devices in the healthcare domain, this review paper is based on the gist of sensing aspects of the micro fabricated paper-based analytical devices (μPADs). The article discusses not only the various used design methodologies and fabrication approaches, but also elucidates the recently reported surface modification strategies, detection mechanisms viz., colorimetric, electrochemical, fluorescence, electrochemiluminescence etc., and sensing characteristics of μPADs along with reported works equipped with the IoT platform for digital readout. In a nutshell, this article summarizes the state-of-the-art research work carried out over the nano functionalized paper-based analytical devices and associated challenges/solutions in the point of care testing domain.

Keywords: Sensor; IoT, μPADs; Sensitivity; Point of care testing; Selectivity

- INTRODUCTION

Initial screening of diseases at the preliminary level is still a challenge across the globe and in this context, micro fabricated analytical devices on paper substrates have been visualized as one of the sustainable and commercially viable solutions in the point of care testing (POCT) domain due to their portability, economy, easy availability, and transportability. Paper-based sensors have recently experienced great progression in terms of their properties and are fulfilling the ASSURED law: viz.(1) Affordable (2) Sensitive (3) Specific (4) Users friendly (5) Rapid and Robust (6) Equipment Free (7) Deliverable to all end users [1][2]. In aforementioned specifications, sensitivity plays a crucial role in the quality of the devices. Analytical devices aimed for POCT can be improved via basic modifications in conventional detection techniques for high-performance sensitivity. The available conservative detection techniques are electrochemical detection, colorimetric detection, chemiluminescence, fluorescence, mass spectrum, and surface-enhanced Raman spectroscopy etc. [3]. The device sensitivity can also be tuned by materials engineering, surface modification, change in design, and fabrication technique. As an example, the detection process can be enhanced by incorporating active catalysts such as silver nanoparticles (Ag NPs), gold nanoparticles (Au NPs), metal oxide nanoparticles, bi and tri-metallic nanoparticles, hyper-branched polymers, graphene and its derivatives, nanoparticles incorporated carbon-based materials, quantum dots, up-conversion nanoparticles etc. Enhancement in sensor’s surface and detection affects numerous technical and viable specifications of the device. These specifications are LOD, specificity, assay speed, detection format, sample matrices and packaging, labeling, stability requirements, and the target cost. The detection sensitivity is often calculated by dividing the slope of the linear range curve by the surface area using the projection area. The different methods designed for different analytes, vary in size, function, and LOD. The functionality of μPADs is broad and can be used for various purposes, for on-site initial screening of human health, water, food, and air. With minute modification in surface of μPADs, not only various diseases have been reported to be diagnosed but also, heavy metal ions in air and water, toxic gases/chemicals in air/water, food/beverages quality etc. The quantification of analytes improved through the integration of electrochemical sensors or strip readers, as well as low detection limit and higher sensitivity achieved through these readers and sensors. Further novel and multi detection of  analytes, improvement in analytical performance, stability and higher shelf life can be achieved through the integration of nanomaterial’s in paper based surface [4]. The sustainability of the lateral flow biosensors (LFB), which is a kind of paper-based biosensor, can be enhanced through the integration of new nano materials, methods, and strategies. The higher sensitivity and lowers detection limit can be achieved through the (1) nanoparticles, which are stable, easy to detect, and provides intense signals (2) developing strategies, which increases signaling quantity per unit of analytes (3) combination of different detection methods, which removes drawbacks of individual detection system and provides high quantification and strong signal in detection zone[5].There is lot of advancement in the microfluidic system for emergency analysis of environmental (soil, air, ecology, river water) [6], food borne pathogens [7][8]. Apart from modifications in design, and surface property enhancement of μPADs, reporting system also plays a crucial role in sensing mechanism of the POCT system. Reporting system basically depends on the sensing mechanism of the μPADs. It can be numeric, image, digital data, graph etc. After the assays it is important to read out the information generated through the optical or electrochemical measurement. It is a significant step for equipment dependent quantitative readouts. But, as per the requirements semi-quantitative and qualitative read outs can be performed. The information gathered can be utilized through a standard analytical-chemical instrumental technique viz., diffuse reflectance spectroscopy, IT communication equipment or electronic equipments such as cell phones, scanners etc.[9]. Smartphones now a days are ubiquitous and are reported to play an important role in performing POCT in sustainable and economical fashion. They not only persist analytical reading capabilities through the screen which can act as a display as well as controller, but also a potential source of capturing the signals via inbuilt camera, audio jack, power port, and ambient light sensors. They are well known medium of a huge storage (device and cloud storage) for data collection and its analysis capability through various software, as well as data transfer ability through wire or wireless (Bluetooth, NFC, Wi-Fi etc.) system. These devices act as a part of internet of things (IoT), through internet connection colorimetric images/ electronic signals can be compared with data samples available in data base and an accurate data interpretation will be available in no time. As well as machine learning (ML) and artificial intelligence (AI) can predict the data via comparing previous data available in data base, which is helpful for analysis of human health, water, food/ beverages, and environmental conditions. Regarding the signal measurement, data transfer, and data analysis, fifth generation (5G) smart-phones have very high capabilities. It can detect visual stimulation at high sensitivity, and image resolution through a high definition (HD) camera, as well as fast data interpretation through high internet and fast processor.

Paper based sensing devices fulfilled all the criteria of ASSURED law and provides sustainable solutions. It is economical, environment friendly, and easy to operate at end users. It is easily applicable for all the diagnostic assays by integrating electrical and optical sensing modalities. It had been demonstrated that it can detect and isolate various bio-targets and pathogens in different matrices including plasma, whole blood, and peritoneal dialysis fluid [10]. An android app is developed for reading the information generated during testing of blood type in paper-based sensor. The barcode like design was developed for easier analysis of length-based information through the smart phone, and it is reported that there is less environment effects in comparison to the electrical and colorimetric signals. The results generated was saved as electronic data and transfer through the text message to the patient/professionals [11].

Figure1 Role of Technology in point of care testing

Further, Moazeni et al. [12] developed a novel peptide modified plastic-paper μPADs for the detection of alpha-fetoprotein (AFP) in human serum. The upper surface of chromatography paper modified with the silver-graphene electrode (20% Ag) and diphenylalanine used in detection zone to increase the sensitivity and stability of the immobilized antibody. That showed detection limit 1 and 10 ngmL-1 in Phosphate Buffer Saline (PBS) and plasma respectively. Adkins et al. [13] developed a method for recognizing fecal indicator bacteria (FIB) through electrochemical and colorimetric detection, two different methods utilize to overcome the disadvantage of a particular method. It has resulted that the detection limit of colorimetric method is higher in comparison to the electrochemical method while the detection time was same.

2. Design layouts and surface modifications in paper based analytical devices

As per the research in the last few years, it is observed that nanomaterials play a vital role in enhancing the electrode surface and bio sensing paper device performance by enhancing analytical signal with the improved available superficial area of μPADs. The modification of paper surface addresses the problems associated with the lack of homogeneity on color measurements that compromise the sensitivity and detectability in clinical diagnosis. Fig.1 demonstrates the modules explored in paper-based devices for detection.

Figure 2 Design modulus explored in paper based deices for detection[14].

2.1- Design Layouts

Fluid in paper-based substrate is transported through the capillary action and it is due to the relationship between adhesive and cohesive force, transportation of fluid in the paper which can be better explained through Lucas-Washburn equation. It describes the one-dimensional capillary flow in a parallel cylindrical tubes bundle [15]. The relation between wicking time and wetted length can be explained through eq. 1 for ideal condition, whereas  simultaneous flow rate in the paper can be expressed as eq. 2 [16], where, l(t) is the distance penetrated by a liquid flowing under capillary pressure into a channel, γ is surface tension of the liquid, r is average pore radius also called as effective pore radius, θ is contact angle between the fluid and boundary wall,  μ is viscosity and t is time, Q represents the volumetric ﬂow rate, k is the permeability of the paper to the fluid, A is the cross sectional area of the channel perpendicular to the flow,  A = W*H (W= width, H = Height), ∆P is the pressure difference along the direction of flow over the length L being characteristic length scale.  The volumetric fluid flow rate varies with respect to the available number of sections and geometry available in a paper-based scheme and can be expressed as eq. 3. But for application purposes, paper-based schemes are used in ambient circumstances where evaporation affects the transportation of fluid, so the (eq.1) can be better transformed to the eq. 4 [17]. The modified form of eq. 1 reported for fluid transportation in paper-based scheme for non-uniform cross-section by predicting filling time t for an arbitrary cross-section {A(l)} and wetted distance {l(t)}[18], can be expressed through eq. 5. The reported effect on contact angle variation due to the hydrophobic barriers explained through modification in eq. 1 and can be expressed through the eq. 6[19].

l(t)=rcosθ2μt                                                        (1)

Q=-kWHμL∆P                                                         (2)

Q  =-∆PinμLikWiHi                                                             (3)

lt=rhcosθ4μq01-e-2q0ht                                      (4)

Dt =kR0μ0lAl'dl'+ 0l[A(l')0l'dxA(x)]dl'          (5)

l(t)=k(1+βd13wcosθbcosθ)σμt                                 (6) Where, Li is the length of the segment i of the channel in the direction of flow, n is the number of connected straight channels, k represents the permeability of the paper to the fluid, ø is porosity of the paper, h is the thickness of the porous channel, q0 is the rate of evaporation and D (=k∆p/μ) is the diffusive coefficient.

Nguyen et al. [20] proposed that the performance and sensitivity of the μPADs could be increased by a simple change in the representative structure of μPADs such as enlarging the detection zone, adding a waste zone, and using an elution step, that can reduce sample loss. The optimized μPADs showed 28 % increase in sensitivity and a 78 % increase in maximum signal for Ni (Ⅱ) samples and a 94 % increase in the maximum signal for Mn (ll). Yang et al.[21] designed five different structures of the paper microfluidic chip to find image grey level statistics for chromogenic reaction and designed Y-Channel, three-channel, rotary array, oblique double channel, and multilayer solid microfluidic chip and concluded that no significant relationship between the color reaction uniformity and the number of channels but found a positive correlation with the reaction contact area. In other work, Li et al. [22] used origami paper-based analytical devices to detect glucose through the three-electrode system and the electrode prepared using direct drawing of a pencil. It (origami paper) showed excellent analytical performance for glucose sensing, reproducibility, and selectivity for common interferents in physiological fluids and exhibited LOD and linear range of 0.05 mM and 1-12 mM, respectively. Park et al. [23] modified a 3D paper-based microfluidic device by utilizing a pressed paper for the detection of c-reactive protein in vertical flow multistep assay, which produced delayed flow and programmed delivery order of the reagents and detected highly sensitive c-reactive protein within  15 min with increased sensitivity and the upper limit of the assays, shows that vertical flow assays with multistep reaction would be much beneficial than lateral flow assays for comprehensive range detection of target analytes. Zhou et al. [24] used 3D rotary μPADs containing three layers paper-based chip and its component, the top layer rotates at 900 to facilitate the detection window on the top layer corresponded to the paper@ZnSequantumdots@ion-imprinted polymer sensing sites of the bottom layer and the design also improved the portability of the device by transferring the liquid phase of ZnSeQDs@ion imprinted polymers to solid glass fiber paper with good sensitivity, selectivity and multiplexed detection. Zhu et al. [25] detected glucose with the help of a tree-shaped paper strip, which showed self-calibration on the test strip and assured the uniformity of the micro flows in different branches at the same time, which helped to collect multiple data of the samples without any delay because of the minimum effect of environmental conditions (temperature, humidity, pressure, etc.) on self-calibration. At the same time, a Y-shaped strip can semi-quantitatively judge whether the concentration in blood serum is low level, average level, or high level, and directly provide a result of "Yes or No."

  - Surface modification via nano functionalization

The inclusive aim of surface modification is to get reliable, quick, and economical testing results and it is required to develop POCT device with better performance, usability, and portability. This can be achieved through the surface medication of the electrodes and detection zones by the smart material. These smart materials are designed materials and play a vital role by offering the capability to capture and release specific components at different conditions by isolating the analyte during sample pre-treatment[26]. The electrodes and the detection zone of the μPADs modified through the various metals’ nanoparticles and bio-polymers in different concentrations. Due to its high surface area these molecules have effective surface for bio-molecule immobilization, better confirmation, and superior biological activities resulting in improved sensing characteristics.

2.2.1 Enhancement through silver nano-particles

The Ag NPs belong to the most substantial conductive Nano-material family due to its small molecular diameter, large surface area, and fast charge transfer, it improves detection capabilities of the device [27][28]. The conductivity of Ag particles increases through the high density of Ag NPs, which can be achieved through small particle and air stability results of Ag NPs; these properties promote the use of Ag NPs as conductive ink in paper-based flexible electronics [29]. It is reported that silver nanoplates behave as an extremely sensitive material towards the inorganic anions [30]. Ding et al. [31] used silver/silver chloride (Ag/AgCl) as a reference element for the ion determination in clinically relevant samples and environmental samples. Ferreira et al. [32] introduced novel optical silver nanoparticles-based sensor for the quantification of ascorbic acid. The sensitivity of the colorimetric measurement of ascorbic acid increases due to the cluster formation growth property of the silver nanoparticles. Ratnarathorn et al. [33] demonstrated the use of Ag NPs for the colorimetric sensing of copper ion (Cu2+) in the paper-based device for the rapid, easy and, economical POCT. Color change was observed through the naked eye due to the tremendously elevated extinction coefficient of the Ag NPs.

2.2.2 Enhancement through nano-bio-polymer

Bio-functionalization is performed for the enhancement of the working substrate. Chitosan is a common bio-functionalized material and its derivatives have been commonly used for the stability of the antibodies due to its cationic properties. It also persists some important properties which makes it exceptional from other biopolymer viz. biocompatibility, easy solubilizes in an aqueous medium, bio-degradability, hydrophilicity etc., due to these exceptional properties chitosan became a great immobilization matrix for the fabrication of a wide range of biosensors [34]. Parween et al.[35] evaluated different bio-functionalization methods of chitosan on paper based schemes for better stability and immobilization of antibodies and found that chitosan modified schemes show better stability at room temperature. On the other hand, sodium hydroxide cross-linked chitosan (CHI-NaOH) bio-functionalized schemes showed stability more than 100 days due to the increment in degree of carboxymethylation and deacetyalation of chitosan, which is maximum in comparison to other functionalized methods. Chitosan is a biopolymer that is extensively used to modify the bio sensing surface and improves capability of the direct electron transfer between a reactive surface and enzyme, by providing a suitable microenvironment. Chitosan has very high surface area supporting cellulose fiber to absorb enzymes and chromogenic substrate [36][37].

2.2.3 Enhancement through gold nano-particles

Gold nanoparticles (Au NPs) are unproblematic to be conjugated with toxins, hormones, immunoglobulins, antibiotics and other biological macromolecules through electrostatic interaction without affecting their biological activity. It is biocompatible, carries size tunable properties and provides a range of color code based on the nano particle size. It is reported that decoration of Au NPs on the surface of rGO-TEPA, enhanced biocompatibility and stability of the rGO-TEPA/Au Nano composite which enhanced the acceleration of the electron transfer to the screen printed electrodes (SPEs) as well as capturing antibodies[38]. Further, it is reported that enzyme modified Au NPs shows high color intensity and sensitivity in comparison to traditional Au NPs [39]. Further, two new optical assay technique direct reduction and bienzymatic methods proposed for the detection of sugars in beverages. The direct reduction of Au NPs provides good sensitivity, while bienzymatic method provides lower cost and simplification, and it does not suffer bleaching due to the high stability of the Au NPs [40]. It is reported that closed Au-bipolar electrode persist excellent conductivity in comparison of graphite electrode, a good choice over the three-electrode system, because it can distinguish the anode and cathode in to separate compartments, avoid any interference and have high current efficiency. Au-Pd used as a catalyst which could catalyze the reduction of hydrogen peroxide (H2O2) for signal amplification.  [41] It showed good sensitivity and selectivity. Due to the following properties, it is highly reported nanomaterials.

Figure 3 Cycle representation of ECL biosensor for the detection of miRNA-155 based on a paper-based BPE with immobilization, displacement reaction,  cycle of target, the immobilization of S1–AuPd.{Reproduced with kind permission} [41].

2.2.4 Enhancement through metal oxide nanoparticles

Contribution of metal oxide nanoparticles (MONPs) is well known, and therefore it has been heavily explored as heterogeneous catalysis. To prepare and understand the catalytic materials, the sol-gel method is utilized due to its exceptional versatility. MONPs show its good biocompatibility, chemical stability, high surface area, and the ability to transfer electrons at a higher speed. These characteristics of MONPs make it perfect for immobilization matrices and transduction platforms. Recent research shows that the modification of electrodes through metal oxides can improve the sensitivity of the electrodes. Nano phase metal oxides with unique optical, electrochemical properties, along with the required surface functionalities and electronic properties, offer fascinating bio-platforms for the interaction of bio-recognition elements with transducers for better signal intensification. Figueredo et al. [42] described the surface modification of μPAD through Fe3O4 NPs, MWCNT, and Graphene Oxide for the detection of glucose. The reported procedures solve the problem of enzymatic activity and give an enhanced analytical performance at low concentration. Evans et al. [43] fabricated a μPAD through CO2 laser engraver, and for the enhancement of analytical sensitivity of the μPAD it immersed in to the suspension of the APTES and modified silica NPS. These NPs can be conveniently entombed in the cellulose structure, which provides strong support for the immobilization of the enzymes. The reported qualitative analysis of lactate, glutamate, and glucose in artificial urine shows improved color intensity and uniformity. Li et al. [44] used a carbon ink working electrode decorated with zinc oxide Nano-wires (ZnO NWs) for the glucose detection with higher sensitivity and lower limit of detection due to high enzyme capturing efficiency of ZnO NWs and high surface to volume ratio. The synthesis of the ZnO NWs was carried out by using hydrothermal route and did not require any light-sensitive electron mediator which explained enhanced bio-sensing stability, and showed sensitivity up to 8.24μA mM-1 cm-2. Sun et al. [45] used CeO2-Au@glucoseoxidase for signal amplification in electrochemical analyses to detect miR-21.

Figure 4 Illustration of the stepwise biosensor fabrication. Reproduced with kind permission from  [45].

2.2.5 Enhancement through bi- and trimetallic Nanoparticles

Bi-metallic nanoparticles of Au, Pd, Ag, Pt show good intrinsic peroxidase-like activity, improve signal amplification [46], and enhanced catalytic activity and selectivity [47]. Nano cluster of different metals have discrete energy levels for sub-nanometer-size centers, it connects the missing link between the particles and metal atoms. They persist many diverse properties viz. high biocompatibility, easy preparation, good photo stability, and low toxicity [48][49]. Zhu et al. [50] used Cu3(PO4)2 nano crystals in the detection zone with GOx and horseradish peroxidase (HRP) to detect glucose in blood serum. It exhibited a flower-like structure that preserves the activity and enhances the stability of enzymes, and allows for the co-immobilization of the enzymes. Pebdeni et al.[51] fabricated a μPAD for the detection of bacterial cells through the colorimetric detection method and resulted that dynamic linear range increases and detection limit decrease for a wide variation of the concentration of S.aureus, while using Bi-metallic Nano cluster of Au/Pt. Wu et al.[52] developed a highly sensitive and specific μPAD for the detection of living cancer cells with high reproducibility, low detection limit, stability, and broad linear range. The paper-based working electrode developed through the Au@Pd, which shows excellent electrochemical property for cyto-sensing and for the good ECL labeling Pt-Ni alloy used with high loading of carbon dots. Arshad et al. [53] developed a simple, selective and sensitive μPAD for the detection of explosive nitroaromatics through colorimetric detection technique. The μPAD functionalized through Au@Ag core-shell nanoparticles with β-crystamine which shows selective determination of tri-nitro toluene (TNT) due to the charge transfer from electron rich β-crystamine to TNT.

Figure 5 Simple mechanisms for the formation of Au/Pt NCs, and detection of S.aureus through paper. Reproduced with kind permission from [51].

2.2.6 Enhancement through nanoparticles combined with carbon-based nanomaterials

Graphene is a single layer of carbon atoms closely packed into a 2-D honeycomb structure arrangement. It shows excellent electrical mobility and electrical conductivity for the fabrication of graphene-based electronic material. It has a small band-gap that greatly contributes in electron conduction from target molecules in electrochemical biosensors. Graphene can provide ultra-high surface area for loading of bio-receptorn. It possesses high sensitivity due to low electronic noise from thermal effects. Deng et al. [54] synthesized graphene oxide@SiO2@CeO2 hybrid nano sheets (GSCs) through wet chemical strategy as an alternative to the commonly employed peroxidase, and found that these Nano sheets possessed high intrinsic peroxidase activity, and high potential for colorimetric indicator. Fan et al. [55] used amino functional graphene, thionine and gold nanoparticles (NH2-G/Thi/AuNPs) for modification of μPADs, and found that it showed a wide linear range and low detection limit due to the NH2-G. The NH2-G exhibits excellent biocompatibility and conductivity that accelerate electron transfer and current signal amplification. Wei et al. [56] used graphene Nano composite (AuNPs/rGO/THI) coated working electrode for the immobilization of DNA aptamer. The THI were utilized to transduce the biological recognition between DNA aptamer and prostate specific antigen, which also served as an electrochemical mediator, AuNPs served as an excellent conductor (Au NPs boost the catalytic activity), and rGO facilitates high electron transfer that showed low LOD, and wide linear range. In another work, dopamine detected by using graphene and sodium dodecyl sulfate surfactant dopamine [57]. Liu et al. [58] used water-dispersible molybdenum oxide quantum dots for highly sensitive and stable sensor to detect H2O2 released from P12 cells. Molybdenum oxide quantum dots was prepared using a facile ultrasonic-assisted hydrothermal method and starting material was MoS2 powder, further synthesized process as shown in the chemical reaction 1, 2, 3, 4.

Zhou et al.[24] used ZnSe quantum dots (lesser toxic than CdTe) for fluorescence detection of cadmium and lead ions. Cincotto et al.[59] used graphene quantum dots for direct oxidation of uric acid, whereas graphene quantum dots with creatininase and a ruthenium serves as electrochemical mediator for creatinine oxidation, and the background current increased significantly due to the larger electro active area of graphene quantum dots at the electrode surface.

2.2.7 Enhancement through up conversion Nanoparticles (UCNPs)

The Up-conversion nanoparticles can be defined as a type of luminescent nanomaterials doped with rare earth elements and can convert near-infrared (NIR) excitation light into visible light ultraviolet emitting light [60]. It has poor water dispersibility and solubility; it is necessary for chemical modification on the surface. UCNPs persist narrow emission spectrum, low toxicity anti-interference from biological auto fluorescence, and anti-stokes shift compared to quantum dots, which are ideal fluorescent labels in FLFIAs [61].  Chen et al.[62] used a paper supported aptasensor for detection of exosomes (cancer diagnostic marker) by using luminescence resonance energy transfer with up conversion nanoparticles with gold nanoparticles and obtained colored image to quantify the luminescence with a range of 1× 104-1 × 108particles/μL and detection limit 1.1× 108 particles/μL. Jiang et al.[63] employed a paper supported aptasensor for the detection of IgE in both buffer solution and human serum with linear range of 0.5-80 ng/mL by using luminescence resonance energy transfer with carbon nanoparticles as energy acceptor and up conversion nanoparticles used as an energy donor.

3. Various sensing mechanisms explored in IoT assisted paper devices

The sensing mechanism for μPADs is characterized in the following sensing techniques: colorimetric, electrochemical, fluorescent, photo electrochemical, chemiluminescence (CL), electrochemiluminescence (ECL), nanoparticles based sensing, spectrometry sensing [64]. Table1 and table 2 summarize the recent progresses carried out on different sensing mechanism used for diagnosis of health and IoT enabled paper based analytical devices respectively.

3.1 Colorimetric sensing

The colorimetric signals achieved on the paper substrate by the reaction of enzymes and detection substance, initial color signals can be observed through the naked eye, but the intensity of the color signals diagnosed through the image sensors, for instance, charge-coupled device (CCD) or complementary metal-oxide-semiconductor (CMOS) camera with high sensitivity. A simple, rapid, sensitive, and cost-effective hybrid μPAD developed for the screening of chlorpyrifos-oxon in human serum. A smart-phone reader used for getting standardized optical detection and software utilized for extracting RGB intensity [65]. Chen et al. [66] used N-ethyl-N (3-sulfopropyl)-3-methyl-aniline sodium salt (TOPS)/4AAP in μPADs for glucose detection, and DHBS/4AAP for uric acid detection, to enhance the sensitivity of the μPADs very effectively, a drying method was employed to reduce the background signal by accelerating the volatilization rate in an incubator resulted in enhanced sensitivity. In another work in this context, Further Li et al. [37] reported a double-layered structure combined with 3D μPAD for the detection of multiple bio-molecules viz. uric acid, glucose, lactate and choline simultaneously. The detection zones of the μPAD were modified through chitosan, which provides a better microenvironment for the enzyme reactivity as well as better adsorption of the chromogenic substrates and enzymes over cellulose fibers. Reported double layer structure showed enhanced sensitivity, colorimetric performance and reproducibility by providing controlled diffusion of colorimetric reagents with in a detection zone, further the detection range and sensitivity increases due to the use of smart-phones as a detector. Jayawardane et al. [67] proposed paper (filter paper) based analytical device for determination of ammonia in wastewater through gas diffusion. This allowed the quantitative conversion of the ammonium ion to molecular ammonia which diffused across the hydrophobic micro porous Teflon membrane of the device into an adjacent hydrophilic reagent zone containing the acid-base indicator 3-nitrophenol or bromothymol blue. Tseng et al.[68] proposed a device based on Jaffe's reaction for performing creatinine detection in the whole blood sample, the reaction zone of 3D μPAD doped with NaOH and picric acid reagent Jaffe reaction induced between the reagent and blood plasma creatinine and concentration detected through the colorimetric method. In another work Tian et al. [69] described a novel method for the molecular recognition and amplification in signals by integrating cross-linked target responsive hydrogel (glucoamylase-trapped aptamer) with the cascaded enzymatic reaction for the μPAD POCT. The proposed device tested for the urine and cocaine in the buffer. Son et al. [70] developed a polydiacetylene-based paper sensor for naked-eye detection of pandemic influenza a (pH1N1) virus. It (PDA-paper chips) changes color under various external conditions (temperature, pH) form blue to red. Further, an app was developed for the detection of viruses at low concentrations. Im et al. [71] developed an optical biosensor for the detection of glucose and lactate as a biomarker for monitoring of cell growth by using a smart-phone with high accuracy and reproducibility with linear range 0.3 – 8.0 mM and LOD 0.3 mM for glucose detection and linear range 0.02 – 0.50 mM and LOD 0.02 mM for lactate detection. Chun et al. [72] developed a μPAD containing both enzymes mediated assay function and sample transporting function for easy and accurate glucose bio sensing and used smartphones as a potential signal transducer. Li et al. [73] constructed a μPAD through the direct synthesis of enzyme-inorganic hybrid nanomaterial’s on the paper matrix, for the in situ growth of GOx@Mn3(PO4)2 hybrid function materials in an inorganic manner solution(MnSO4 and KH2PO4) containing a diluted enzyme glucose oxidase and GOx pipette on to cellulose paper. The centrifugation and the dry process eliminated through this new approach of the in-situ growth of an enzyme-inorganic hybrid on a paper matrix. Calabria et al. [74] developed a smartphone biosensor in which reagents are necessary for analysis co-entrapped on paper in a wafer-like bilayer film of polyelectrolytes to detect L-lactate in oral fluid with 0.1mML-1 limit of detection. A simple, fast and sensitive approach for detecting Zika virus executed with a smartphone and μPADs, an image of the detection zone is captured through the smartphone camera for real-time quantification. The color intensities are evaluated through the software [75].  In another work, smart-phone is used for the colorimetric detection of Hg2+ in an inkjet-printed paper-based sensor modified with silver nanoparticles. The quantitative analysis performed through calculating intensity of Ag NPs  through color detector app and smart-phone [76]. In another work, Cu2+ ion presence was determined in sugar cane spirits through the smart-phone camera, by functionalizing paper-based device by the cuprizone. The proposed method is simple, precise and fast in terms of qualitative and quantitative determination [77].

Figure 6 illustrates (1) smart phone based device with assays of glucose and lactate based on reaction time [71], (2) Illustrate main component of portable detection system and detection box with effect of reaction temperature(a) and reaction time(b) on green + blue intensity [68].

3.2 Electrochemical sensing

Electrochemical biosensors can be fabricated in different ways, including microfluidic paper-based analytical devices (μPADs), microelectrode array (MEA), and single large electrode. The most common electrochemical detection methods include cyclic voltammetry, chronoamperometry, and differential pulse voltammetry (DPV) measurements. DPV possesses higher detection precision than chronoamperometry and cyclic voltammetry by reducing background current [78] and can be used for detection of different POCT Hemoglobin[79], Glucose, etc. in food and beverages. Cao et al. [80] constructed a novel 3D microfluidic electrochemical glucose biosensor, where reduced graphene oxide tetraethylenepentamine (rGO-TEPA/PB) used as a working electrode. The prepared working electrodes showed high conductivity and good electro catalytic reduction towards H2O2 due to large specific area. The prepared biosensor showed broad linear range and low detection limit while detecting glucose in sweat and blood. In another work, Cao et al.[81] designed and constructed a novel μPAD for the determination of human chorionic gonadotropin (HCG). Immuno-filtration techniques were introduced for the fabrication of the device, and periodate oxidation was used for the covalent conjugation of the antibodies on the hydrophilic zone of the SPEs. This showed pleasing specificity and sensitivity. Maattanen et al. [82] developed a low-cost paper-based platform for electrochemical analyses. Working and counter electrodes directly printed gold stripes and reference electrode printed with silver stripes onto which silver chloride layer was deposited electrochemically showed performance equal to the conventional electrode with good pH sensitivity even after five weeks of storage. Ruecha et al. [83] prepared a Nano composite of graphene /  polyvinylpyrrolidone / polyaniline (G/PVP/PANI) and used it for modification of paper-based biosensors via electro-spraying. Adding a small amount of PVP, improved the dispersion of graphene in Nano composites, which increased electrode conductivity and sensitivity of the biosensor, prepared biosensors provide a cost effective, portable and disposable paper-based device.  Wang et al.[84] detected carcinoembryonic antigen (CEA) and NSE for identifying lung cancer by fabricating a paper-based device via wax printing and screen printing with functions of sample auto-injection and sample filtration with the help of NG-THI-AuNPs and prussian blue- poly(3,4-ethylenedioxythiophene)-gold nanoparticles (PB-PEDOT-AuNP), further synthesis of PB-PEDOT-AuNPs Nano composite explained through the chemical reaction 5 and 6.

Fe3+ + EDOT                              Fe2+ + PEDOT    (5)

Fe2+  + [Fe(CN)6]3-                 Fe3[Fe(CN)6]2 (PB)   (6)

The modified electrode worked for promoting the electrode transfer rate as well as immobilization of the CEA, and NSE aptamers. Noiphunget al.[85] used Prussian blue modified electrode to detect glucose in whole blood samples using dumbbell-shaped µPADs. Ruiz-Vega et al.[86] used magnetic beads to develop electrochemical magneto-immunosensor that performed electrochemical detection and washing on the chip with minimum intervention of the user to detect matrix metalloproteinase 9 (MMP-9). It stated that it was a single step magneto immunoassay for MMP-9 carried out through 8 screens printed electrode and a movable multiplexed fluidic module. The electrochemical detection, magnetic bead confinement under the static condition and reagent adsorption, magnetic bead washing underflow condition provided by the paper device, showed fast simple, and sensitive assay formats. Wei et al. [56] used graphene Nano composite for the detection of PSA. The modified electrode (AuNPs/rGO/THI) showed excellent conductivity and gave a specific peak current associated with an electrochemical mediator of THI. Sun et al.[45] detected miR-21through synthesized Au Nano rods by In-situ growth methods in μPADs for the enhancement of conductivity and catalyzed glucose further electro-catalyzed by CeO2 and gives a wide linear range and low detection limit. Manbohiet al. [57] detected dopamine in blood and urine samples to enhance the sensitivity and selectivity of μPAD toward dopamine. MWCNT, graphene, and Fe2O4 nanoparticles were utilized and sodium dodecyl sulfate (SDS) were selected for efficient detection of dopamine. Wang et al. [87] developed a highly sensitive label-free μPAD for detecting CEA using (NH2-G)/Thi/Au NPs Nano composites (synthesized and coated on the screen-printed electrode) to enhance detection sensitivity. Cincotto et al. [59] developed a device for detecting two analytes at two different spots with the help of one working electrode, through surface modification with graphene quantum dots (QDs) for uric acid oxidation, whereas graphene QDs, creatininase enzyme, and a ruthenium electrochemical mediator for creatinine oxidation. It showed a low detection limit 8.4 nmL-1 for uric acid and 3.7 nmL-1 for creatinine with high sensitivity, selectivity, and reproducibility. The electro catalytic mechanism at the surface of electrode can be described as chemical reaction 7, 8, and 9.

Chaiyoet al. [88] introduced a novel μPAD by modifying a screen-printed carbon electrode with cobalt phthalocyanine, graphene, and an ionic liquid for the non-enzymatic detection of glucose that showed fast electron transfer kinetics and excellent conductivity. It comprised a detection limit of 0.67 μM and a wide linear range 0.01-1.3 mM and 1.3-5.0 mM for glucose. The electro catalytic process can be expressed as chemical reaction 10, 11, and 12.

Adkins et al. [89] used microwire electrode as a substitute for the screen-printed electrode for high sensitivity, low resistance and increase in current density relative to carbon electrode and detected glucose, fructose, and sucrose by using a Cu electrode in alkaline solution with a limit of detection 270 nM, 340 nM, and 430 nM respectively. Ruiz-Vega et al. [90] developed a paper-based double-sided screen-printed electrode for the detection of malaria Plasmodium falciparum lactate dehydrogenase (pfLDH) through a magnetic bead and found that μPADs could be exploited to simplify magneto-immunoassay handling, taking magnetic beads closer to the testing required. Xu et al.[91] demonstrated a POCT system which includes handheld electrochemical analyzer and smartphone, for the detection of heavy metals (Cd2+, Pb+2, Cu+2, Hg+2) with high sensitivity, Installed APP in the smartphone simultaneously controlled the analyzer as well as receive data and further do the action for the measurements and plotting the graphs of the results.  Fan et al. [55] reported a wireless POCT device with μPADs system to detect neuron-specific enolase (NSE) through an electrochemical measurement system to enhance the sensitivity modification done with Nano composites of NH2-G/Thi/AuNPs. The detected results are stored in EEPROM memory automatically. Simultaneously results displayed in android based smartphone via Bluetooth in real-time.

Figure 7 (a) illustrates the Systematic representation of wireless POCT system with μPAD, electrochemical system, and smart phone with customized app operation [55], while figure (b) illustrates systematic layout of glucose detection through μPAD with IOT based analysis [85].

3.3 Fluorescence sensing

Fluorescence detection is generally performed in a closed light environment at a definite wavelength. The enclosed space provides stability, closed light environment, and numerous excitation sources with different specific wavelengths, which makes him more demanding compared to colorimetric detection technique. Jiao et al. [92] constructed a vertical-flow paper-based device for detecting three cancer biomarkers CEA, AFP, and cancer antigen (CA199) for rapid detection within 5 minutes with low sample volume and immune response time, fluorescence image captured through the mobile phone under an ultraviolet lamp. Zhou et al. [24] used 3D rotary μPAD for the detection of  Cd+2 and Pb+2 by using a novel fluorescent ZnSe QDs with ion imprinting technology; it enabled them to realize specific and multi-channel determination of the ions. Guzman et al. [93] developed an integrated platform consisting of μPAD and a portable detection system for detection of low concentration CH2O through Hantzsch reaction (eq.13)

HCHO + AA+ 2(AAA)        Dihydropyridine derivative       (13)

Between ammonium acetate (AA) and acetoacetanilide (AAA) reagent and formaldehyde, the AA/AAA indicator coated on reaction zone of μPAD and resultant fluorescent dihydropyridine derivative observed through CMOS camera with a detection limit 0.2 ppm, analytical curves explained in Fig.5. Ueland et al.[94] developed a novel μPAD to extract, filter, and pre-concentrate explosives from the soil for direct analysis; it involves fluorescence quenching based detection. Liu et al.[95] fabricated a paper-based loop-mediated isothermal amplification (LAMP) sensing chip for the immediate detections of multiple target genes through different pathogenic bacteria, and analysis is done through the IoT for POCT application. The POCT system consists paper sensor and a portable instrument built on IoT platform. Paper sensor provides the functionality of reagent storage, transportation of sample and amplification of nucleic acid. While IoT instrument uses an Arduino microcontroller as fluorescence image collector, temperature controller, and data storage in cloud through a Wi-Fi network. In another work, a dual layer μPAD developed for the detection of ROR1+ cancer cells through whole cell captured imaging and antigen/cell fragment. Where flow fluorescence imaging and velocity analysis done through the smart phone camera [96]. Further, Yetisen et al. [97] demonstrated a paper-based μPAD scheme for the quantitative detection of electrolytes in tears, for early finding of the ocular disorders. Functionalization of detection region done through the fluorescent crown ethers, sensitive for the mono- and divalent electrolytes and smartphone utilized as the readout device for the fluorescent outcomes. Zamora-Galvez et al. [98] used a novel design which is based on lateral flow technology in combination with QDs and the usage of graphene oxide for highly sensitive protein detection(IgG), to enhance the sensitivity of fluorescence μPAD assays QDs printed on paper which successfully showed the limit of detection 1.35 ngmL-1 for standard buffers and 6.30 ngmL-1 for human serums. Anjana et al. [99] used carbon dots doped with sulfur and nitrogen prepared via microwave technique to detect bilirubin in human serum with a detection limit of 0.12 nM. Citric acid and L-cysteine used for the source of carbon and nitrogen, sulfur, respectively. Sutariya et al. [100] used aminopyrine linked calyx[4]arene for highly sensitive and selective detection of As3+, Nd3+, and Br-. In another work, Chen et al.[101] reported a novel naphthofluorescein- based probe ZN-2 which persist long wavelength better sensitivity and faster response in comparison to fluorescein based probe ZN1. This paper-based system showed great potential of POCT system for detecting analytes in real samples.

Figure 8 Part (1) shows a schematic design illustrating the housing, temperature management component, microcontroller, and fluorescence detection optics of an IoT-based LAMP detector[95], Part (2) represents setup for detection of low concentration formaldehyde (a) portable detection set-up (b) components used in detection box[93].

3.4 Chemiluminescence (CL) sensing

It is an outstanding recognition method due to its low cost, wide calibration range, high sensitivity, low cost, simple instrumentation, and low background. In the generalized form signals of CL are generated through the mixing of CL reagent with the oxidizing reagent in the presence of a catalyst (horseradish and metal ions). Li et al. [102] fabricated a double-layer 3D μPAD with high resolution temporally resolved chemiluminescence emissions for multiplexed detection of glucose, cholesterol, lactate, and choline, the detection zone modified with the cobalt ion and different oxidase and obtained linear range for cholesterol (0.01-0.4 mML-1), glucose (0.01-1.0 mML-1), lactate (0.02- 5.0  mML-1), and choline (0.001-1.0 mML-1) with detection limit 8, 15, 6, 0.07 μML-1 for glucose, lactate, choline and cholesterol respectively. Liu et al.  [103] used a novel molecularly imprinted polymer μPAD to detect dichlorvos, the molecularly imprinted polymer (MIP) layer synthesized and adsorbed on the paper for imprinting of dichlorvos. Liu et al. [104] detected dichlorvos (DDV) through paper chromatography, development of DDV was good on paper and separated through the water-soluble vitamins and metal ions. Li et al. [105] detected simultaneously three biomarkers (H-FABP, copeptin, and cTnI) through CL detection method in one sample. Simple 3-D sandwich-type CL immunoassays constructed with different detection zone for individual detection of the biomarkers. The amplified signals were obtained through the dual-signal amplification technique by employing Au NPs and Co(II)-Ab2-luminol-Au NPs for primary and secondary antibody, respectively.

3.5 Electrochemiluminescence sensing

ECL has an advantage over CL, electrochemistry, and fluorescence in ultrasensitive detection and testing using aptamer, nucleic acids, imprinted polymers, and affinity-based sensing elements. Further advances are happening in the field of signal enhancement and forming integrated systems [106]. Delaney et al. [107] described the first approach of electrochemiluminescence at the paper microfluidics combined with screen printed electrodes to detect 2-(dibutylamino) ethanol (DBAE) and nicotinamide adenine dinucleotide (NADH) to the level of 0.9 μM and 72 μM. Further the mechanism involved for the making of ECL by using Ru(bpy)32+ and DBAE is projected through the chemical reaction 14, 15, 16, 17, 18.

Wang et al. [41] used closed Au-bipolar electrode(BPE) for sensitive detection of miRNA-155 and hybridization chain reaction used for the modification of cathode of a bipolar electrode with DNA (S1)-AuPd nanoparticles which catalyzed H2O2 reduction and (Ru(bpy)32+/TPrA) used as an anode of bipolar electrode. It stated that closed Au-bipolar electrode persist excellent conductivity in comparison of graphite electrode, a right choice over the three-electrode system because it can distinguish the anode and cathode into separate compartments, avoid any interference, and have high current efficiency. AuPd used as a catalyst which could catalyze the reduction of H2O2 for signal amplification. Chen et al. [108] developed ECL bio sensing platform for detecting hepatitis B surface antigen (HBsAg) with a modified magnetic suspension sandwich immunoassay method that gave high sensitivity, reliability, and stability. Chen et al. [109] developed a handheld paper-based bipolar electrode- ECL system for detecting glucose in phosphate buffer solution and artificial urine with high sensitivity, stability, and reproducibility through a compact device containing lithium battery for power supply and smartphone for the signal readout. Liu et al. [110] prepared a novel paper fluidic crossing channel closed- bipolar electrode in which multiple band-shaped closed-bipolar electrodes situated perpendicular to two parallel channels and multiples detection detected at reporting channel, and analyzed linear glucose range up to 0.08- 5mM and LOD 0.03 mM, as well as device, showed duplex detection potential for glucose and uric acid, for the particular process reaction mechanism can be described through chemical reaction 19, 20, 21, 22, 23.

Figure 9 A systematic diagram of a hand-held system for Electrochemiluminescence sensing [109]

4. IoT enabled devices

In everyday life IoT based equipment, viz. digital cameras, mobile phones, webcams, scanners, etc., play an essential role in creating life easy. These physical objects are connected with software, sensors, and other technology for the sole purpose of exchanging data with other systems and devices through the internet. Such devices can readily be applied for easy and accurate detection of analytes for point of care detection in μPADs, indicator papers, etc. It is highly advantageous in fact of simplicity, cost, and portability. It provides many opportunities for point of care detection of human health, food/beverages, agriculture, and the environment[111][112].

4.1. Scanners

2-D optical scanners have been used as a low-cost detection system for μPADs. It works on diffuse reflectance spectroscopy principle, and uses filters for binning light into 3 or 4 wavelengths prior to detection by CCDs instead of separating light. The data generated through the phone camera and scanners habitually reproved through the device software, which generally hamper the resurgence of the raw data. This pre-proceeding of the data can lead to the difference between colorimetric read-out, and can hamper the final result. It is suggested that the issue can be alleviate by getting the raw data through the device, in place of using external scanning calibration through a known object. Frequently, the flatbed scanners are preferred for the biomedical application. Because, its quantitative analysis is coupled through the μPADs, which is favorable and economical for the readouts. In flatbed scanners, illumination condition, and the focal distance is constant during entire single signal readout, and also for the consecutive readouts. It generates better images, and can be helpful in reproducibility [111][113].

4.2 Smartphones

Similar to the scanners, smart-phones are also used for the economical quantitative analysis in paper-based devices. These are portable, light-weighted, and economical and no special training required for the operation. The data can be transmitted from direct hardware to the off-site of the specialist, as well as can be directly analyzed through the customized software. It can be utilized for the different analysis methods viz.  colorimetric, fluorescence, electrochemical etc. It can display the data in the form of chart, figure, numeric, and in the form of text message. A wide display, high processor, and a huge storage help to qualitative and quantitative analysis of data. It can directly send the command through the internet which can further analyzed through cloud storage and the patient/ specialist can get the qualitative and quantitative data in no time at the remote location.

The current decade is the witness of the development of μPADs inbuilt with smart-phones and IoT platform which are frequently used for the point of care testing for human health [65] [114] [115] [96] [116] [117] [118] [119], food and beverages [93], water quality [120] [121] [122], and heavy metal ions [123] [124] [77] [76]. Recent progress shows various technical advantages and cost-effectiveness of diagnosing diseases in μPADs, through a smart-phone. Still, it persists multiple physical and technical limitations can be overcome by better-implementing filters, using advanced image processing software, and developing unique platforms [125].

Figure 10 Systematic representation of internet of things in point of care testing.

5. Limitations of µPADs for POCT

μPADs for POCT provides economical solutions for the preliminary level diagnosis. It offers a primary screening tool for healthcare to the end-users. Yet, there exists a lot of scope to improve the performance of these devices because of their varying sensitivity and specificity; it may cause altered results in terms of positive and negative. Due to the portable property of the POCT device, the test strips and reagents may be exposed to environmental conditions (cold, heat, humidity etc.), humidity affects the colorimetric test stripes, altitude affects those tests, which is sensitive to oxygen (viz. glucose and blood gas tests). Disease metabolites, drugs, and diet can affect the POCT test results as compared to laboratory test results[126]. The false-negative result of a home pregnancy test kit causes an increased in pregnancy and sexually transmitted infections [127]. The home pregnancy test kit's false-positive results lead to significant patient anxiety and pointless interventions [128]. It is also observed that the point of care glucose meter's accuracy and clinical performance is hindered at some level of performance [129], through the unsuitable glucose meter may overestimate and underestimate the glucose level, which may hamper the insulin dose of the patient [130]. At-home HIV test kit (OraQuick At-Home HIV test) approved through the US food and drug administration (FDA) shows specificity(99.98%) and sensitivity(92%) that means among the 12 test 1 test result will be false negative [131]. Apart from the fabrication, sensitivity and specificity, handling of the POCT device by the end-users (domestic or medical services)  is the major challenge, which affects the final results of the tests[132]. Therefore it is required to evaluate errors and risk of errors in POCT more carefully, it can be achieved by bearing in mind the whole testing procedure by means of well- designed studies aiming to improve outcomes [133].

6. Possible routes to overcome the limitations of µPADs for POCT

Reported devices have certain limitations when they are being used beyond manufacturers’ recommendations. Errors can be occurred due to the number of operators involved, test specimen volume, test locations. The quality of a μPADs can be estimated through its specificity and sensitivity, which is dependent on the four main functions (1) Detection methods and various substrates (2) Methods of result determination (3) Robustness of testing device (4) Variation in batch production.

Errors in the producing the accurate results on its multiple usage can be reduced by employing risk management during the development of new methods, by documentation of common hazards, limitations, and instruction for use. Risk management examines each and every step of the testing process for finding the weakness of the procedure where errors may occur[126]. The detection methods diverge from optical to electrochemical; for each available detection method, a range of substrate is known for getting the best possible specificity, sensitivity, and cost reduction. The capillary flow speed is a dependent part of the substrate use; as the capillary flow speed increases, the value of specificity increases, and the value of sensitivity decreases simultaneously, directly affecting the test signals. For colorimetric determination, the test result is determined through the naked eye or by the end-user camera. The signal readout is an illumination-dependent term in colorimetric and depends on end-user judgment, mainly when the signal value is near the porch. The robust behavior of the testing devices affects the performance of the devices, the reagents used for the analysis viz. antigens or antibodies and enzymes should endureinsensitivecircumstance during testing, storage and shipping[134]. The testing signals also depend on the atmospheric conditions (temperature, humidity etc.), directly affecting the sample migration speed and stability of enzymes and reagents, shows variation in end signals. While during batch production, the reproducibility of μPADs gets hampered. By keeping these points in mind, the specificity and sensitivity of the μPADs can be further improved at clinical and domestic levels.

7. Conclusion

Diagnosis of disease in low-resource settings is challenging due to the unavailability of trained personal and resources. Diagnosis in many countries is generally performed through histopathology assessment, which needs dedicated space, equipment, professional staffing. Conventional mode to conduct the initial screening of diseases is performed on the medical centers which are expensive, time-consuming, and require trained personals for machinery operation. Thus, for the on-site clinical diagnosis, IOT enabled POCT can be viable solution in the context of portable, economical with reasonable accuracy, and increased sensitivity of the devices. There is no requirement for skilled medical staff and heavy equipment in this case, and it can be performed at home in a remote location. Advance devices including smart phones can play a vital role and be a part of sustainable and economical solution in the enhancement of disease diagnostics. It provides high computing power, high-definition display, and user interface. Along with various facilities, smart-phones also offer a high-performance camera with adequately huge numerical aperture, storage capability, ability to transfer images and analysis. With the use of IoT, sustainable, cost effective as well as environment friendly POCT devices can be developed which will help for the initial diagnosis of the diseases. Herein this review article, we briefly discussed about the application and development of diverse detection platforms of a paper-based microfluidic device for point of care testing, with their design layout, surface modification, sensing mechanism, and performances, with various detection methods. These platforms enabled the transition from qualitative analysis, semi- quantification to quantitative detection in the POCT field and consequently promoted the study and development of the POCT. These platforms were demonstrated to show desirable performance in the detection of microfluidic chips under specific assay conditions. Despite the diversity of detection methods, as well as detection speed, dimensions, study purpose, and analytes of interest, all the detection platforms provided convincing and rapid test results within the allowable error range. It is observed that by modifying the surface of the detection zone and electrodes, accuracy level of the results increases in terms of LOD, linear range, and detection time.

As a concluding remark, this review work summarizes the recently published work on the advanced paper-based analytical devices for point-of-care testing for the initial diagnosis. There is a lot of research scope to develop rapid, secure, sustainable and low-cost detection of the target analytes with greater accuracy. Furthermore, the future research scope lies to find out the sustainable solution in the healthcare domain to explore the device's capability by modifying the design to achieve ultra-sensitivity along with the wireless read out using IoT enabled platform from the end user’s perspective.

Acknowledgments: Authors gratefully acknowledge Start Research Grant (SRG/2020/001895) provided by Science and Engineering Research Board, Department of Science and Technology, India.

Conflicts of Interest: The authors declare no conflict of interest. The funders had no role in the design of the study; in the collection, analyses, or interpretation of data; in the writing of the manuscript, or in the decision to publish the results.

Author Contributions: Conceptualization, VK and AG; resources, AG.; writing—original draft preparation, VK.; writing, review and editing, SP, UTN, KM; supervision and project administration, AG.

Table 1- Paper based health diagnostic devices with various sensing mechanisms.

| Sr. No. | Author (s) | Fabrication Method | Diagnosis | Method | Enhancement | Sample | Range | LOD |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Cao et al.[80] | Photolithography and Screen Printing | Glucose | Electrochemical | Prussian Blue- rGO-TEPA/PB | Blood | 0.1mM-25 mM | 25 μM |
| 2 | Noiphung et al.[85] | Wax Dipping | Glucose | Electrochemical | PB-SPEs | Whole Blood | 0-33.1 mM | - |
| 3 | Li et al.[44] | Wax Printing | Glucose | Electrochemical | ZnO NWs | Human serum | 0-15 mM | 59.5 μM |
| 4 | Chaiyoet al.[88] | Wax Printing | Glucose | Electrochemical | CoPC/G/IL/SPCE | Honey | 0.01-1.3 mM And 1.3-5.0 mM | 0.67 μM |
|  |  |  |  |  |  | White wine |  |  |
|  |  |  |  |  |  | Human Serum |  |  |
| 5 | Deng et al.[54] | Photolithography | Glucose | Colorimetric | GO@SiO2@CeO2 Nano sheets (GSCs) | Human serum and urine samples | 0.5-30 mM | 9 nM |
| 6 | Figueredo et al. [42] | CO2 laser cutting | Glucose | Colorimetric | MNPs | Artificial Urine | 0.05- 1 mM | 43 μM |
|  |  |  |  |  | MWCNT |  | 0.05- 1 mM | 62 μM |
|  |  |  |  |  | GO |  | 0-1 mM | 18 μM |
| 7 | Soniet al.[135] | ---- | Glucose | Colorimetric | - | Saliva | 9-1350 mg/dL | 22.2 mg/dL |
| 8 | Zhu et al.[25] | Normal cutting of chromatography paper | Glucose | Colorimetric | GOD/HRP | Serum | 1-11 mM | 0.3 mM |
| 9 | Zhu et al.[50] | Wax Printing | Glucose | Colorimetric | Cu3(PO4)2 | Serum | 0.1-10.0 mM | 25μM |
| 10 | Gabriel et al.[36] | Stamping | Uric Acid | Colorimetric | Chitosan | Artificial serum/ human tear sample | 0.1-1.0 mM | 0.037 mM |
|  |  |  | Glucose |  |  |  |  | 0.023 mM |
| 11 | Chen et al.[66] | Photolithography | Uric Acid | Colorimetric | Bi-enzyme (GOD/UAO) | Human Serum | 50 μM- 1.0 mM | 43.1 μM |
|  |  |  | Glucose |  |  |  |  | 38.1 μM |
| 12 | Ruecha et al.[83] | Paper-wax printing | Cholesterol | Electrochemical | G/PVP/PANI | Human Serum | 50μM- 10mM | 1 μM |
| 13 | Wang et al.[84] | Wax Printing and Screen Printing | CEA | Label Free Electrochemical | NG-THI-Au NPs | Serum | 0.01-500 ngmL-1 | 2 pgmL-1 |
|  |  |  | NSE |  |  |  | 0.05-500 ngmL-1 | 10 pgmL-1 |
| 14 | Fan et al.[55] | Wax Printing and Screen Printing | NSE- Lung Cancer | Electrochemical | NH2 -G/Thi/AuNPs | Serum | 1-500 ngmL-1 | 10 pgmL-1 |
| 15 | Evans et al[43] | CO2 laser engraver | Lactate | Colorimetric | Silica Nanoparticles | Artificial Urine | 0.63-3.75 mM | 0.63 mM |
|  |  |  | Glucose |  |  |  | 0.5-10 mM | 0.50 mM |
|  |  |  | Glutamate |  |  |  | 0.25-7.50 mM | 0.25 mM |
| 16 | Liet al.[37] | Wax Screen Printing | Glucose | Colorimetric | Chitosan | Serum | 0.01-10 mML-1 | 3 μM |
|  |  |  | Uric acid |  |  |  | 0.01-5.0 mML-1 | 5 μM |
|  |  |  | Lactate |  |  |  | 0.04-10.0 mML-1 | 30 μM |
|  |  |  | Choline |  |  |  | 0.04-24.0 mML-1 | 10 μM |
| 17 | Ruiz-Vega et al.[86] | SPCE | MMP-9 | Electrochemical | Magnetic Beads | Plasma | 0.03-2 ngmL-1 | 0.01 ngmL-1 |
| 18 | Wei et al.[56] | Wax Printing and Screen Printing | PSA | Electrochemical | AuNPs/rGO/THI | Serum | 0.05-200 ngmL-1 | 10 pgmL-1 |
| 19 | Wang et al.[41] | Wax Printing, Screen Printing, In-situ | miRNA-155 | ECL | AuPd NPs |  | 1 pM-10 μM | 0.67 pM |
| 20 | Jiao et al.[92] | Wax Printing | CEA | Fluorescent Immunoassay |  | Serum | 0.1-1000 ngmL-1 | 0.03 ngmL-1 |
|  |  |  | AFP |  |  |  | 0.1-1000 ngmL-1 | 0.05 ngmL-1 |
|  |  |  | CA199 |  |  |  | 0.1-1000 UmL-1 | 0.09 UmL-1 |
| 21 | Chen et al.[108] | SPE | HBsAg | ECL | Magnetic Beads | Serum | 34.2pgmL-1 – 34.2 ngmL-1 | 34.2pgmL1 |
| 22 | Cao et al.[38] | Photolithography | AFP | Electrochemical | rGO-TEPA/Au | Serum | 0.01-100.0 ngmL-1 | 0.005 ngmL-1 |
| 23 | Sun et al. [45] | --- | miR-21 | Electrochemical | CeO2 | Human Serum | - 1000 fM | 0.434 fM |
| 24 | Manbohiet al.[57] | Wax Stamping | Dopamine | Electrochemical | Graphene | Blood and Urine | 0.5-120 μM | 0.01 μM |
| 25 | Cao et al.[81] | Photolithography and Screen -printing technology | HCG | Electrochemical | GNPs | Serum | 1.0mIUmL-1 -100.0 IUmL-1 | 0.36 mIUmL-1 |
| 26 | Wang et al.[87] | Wax Printing and Screen Printing | CEA | Electrochemical | NH2-G/Thi/AuNPs | Serum | 0.05-500 ngmL-1 | 10 pgmL-1 |
| 27 | Cincottoet al.[59] | Craft Printing | Uric Acid | Electrochemical | Graphene Quantum Dots | Human Serum | 0.010- 3.0 μmL-1 | 8.4 nmL-1 |
|  |  |  | Creatininase |  |  |  |  | 3.7 nmL-1 |
| 28 | Zamora-Galvez et al. [98] | Printing | IgG | Fluorescence | QDs-GO | Serum | 1-1000 ngmL-1 | 6.30 ngmL-1 |
| 29 | Anjana et al.[99] | - | Bilirubin | Fluorescence | S,N-CDs | Serum | 0.2-2.0 nM | 0.12 nM |
| 30 | Ruiz-Vega et al.[90] | SPCE | Malaria(pfLDH) | Electrochemical | Magnetic Beads | Whole blood sample | 0.006-1.5 % in culture RBC | 200 ngmL-1 |

Table 2- IoT enabled POCT devices for various diagnosis reported in last 4 years.

| S. No. | Author | Fabrication Method | Diagnostic | Method | LOD | ### Data acquisition and analysis |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Liu et al. [95] | Paper substrate | Nucleic acid | Fluorescence | 2*103 copies/μL | IoT microprocessor for data collection, cloud storage, |
| 2 | Chen et al. [109] | Wax screen printing | Glucose | Electrochemiluminescence | 30 μM | Smartphone imaging, wireless transmission of ECL signal to PC |
| 3 | Tsagkaris et al. [65] | Wax printing | chlorpyrifos-oxon | colorimetric | 0.033 μg L−1 | Smartphone reader for optical detection, smartphone app for providing results |
| 4 | Im et al. [71] | Wax printing | Glucose and lactate | colorimetric | 0.3 and 0.02 mM | Smartphone as optical sensing system, color intensity quantification through software. |
| 5 | Calabriaet al. [74] | Layer by layer | Salivary lactate | colorimetric | 0.1mML-1 | Smartphone for detection, computational online data analysis, for data transfer. |
| 6 | Kaarj et al. [75] | Wax printing | Zika Virus | Fluorescence | 1 copie/μL | Smartphone for image quantification and data analysis. |
| 7 | Monisha et al. [76] | Inkjet printing | Hg+2 | Colorimetric | 10 μg L−1 | Smartphone for image quantification and data analysis. |
| 8 | Franco et al. [77] | Wax Printing | Cu2+ | Colorimetric | 0.034 mgL-1 | Smartphone for image quantification and data analysis. |
| 9 | Ulep et al. [96] | Wax Printing | ROR1+ cancer cells | Fluorescence | 0.1 cells/μL | Smartphone for fluorescence microscopy and image processing and software for analysis. |
| 10 | Kou et al. [114] | Cut and stack | Glucose and uric acid | Colorimetric | - | Smartphone for image quantification and data analysis. |

References

[1]	A. Nilghaz, L. Guan, W. Tan, and W. Shen, “Advances of Paper-Based Microfluidics for Diagnostics - The Original Motivation and Current Status,” ACS Sensors, vol. 1, no. 12, pp. 1382–1393, 2016.

[2]	A. Nilghaz, S. Bagherbaigi, C. L. Lam, S. M. Mousavi, E. P. Cόrcoles, and D. H. B. Wicaksono, “Multiple semi-quantitative colorimetric assays in compact embeddable microfluidic cloth-based analytical device (μCAD) for effective point-of-care diagnostic,” Microfluid. Nanofluidics, vol. 19, no. 2, pp. 317–333, 2015.

[3]	S. C. B. Gopinath, T. Lakshmipriya, Y. Chen, W. M. Phang, and U. Hashim, “Aptamer-based ‘point-of-care testing,’” Biotechnol. Adv., vol. 34, no. 3, pp. 198–208, 2016.

[4]	C. Parolo and A. Merkoçi, “Paper-based nanobiosensors for diagnostics,” Chem. Soc. Rev., vol. 42, no. 2, pp. 450–457, Dec. 2013.

[5]	D. Quesada-González and A. Merkoçi, “Nanoparticle-based lateral flow biosensors,” Biosensors and Bioelectronics, vol. 73. Elsevier Ltd, pp. 47–63, 05-Nov-2015.

[6]	C. Te Kung, C. Y. Hou, Y. N. Wang, and L. M. Fu, “Microfluidic paper-based analytical devices for environmental analysis of soil, air, ecology and river water,” Sensors Actuators, B Chem., vol. 301, no. August, p. 126855, Dec. 2019.

[7]	B. Al Mughairy and H. A. J. Al-Lawati, “Recent analytical advancements in microfluidics using chemiluminescence detection systems for food analysis,” TrAC - Trends Anal. Chem., vol. 124, p. 115802, 2020.

[8]	M. Puiu and C. Bala, “Microfluidics-integrated biosensing platforms as emergency tools for on-site field detection of foodborne pathogens,” TrAC - Trends Anal. Chem., vol. 125, p. 115831, 2020.

[9]	D. Quesada-González and A. Merkoçi, “Mobile phone-based biosensing: An emerging ‘diagnostic and communication’ technology,” Biosensors and Bioelectronics, vol. 92. Elsevier Ltd, pp. 549–562, 15-Jun-2017.

[10]	H. Shafiee et al., “Paper and flexible substrates as materials for biosensing platforms to detect multiple biotargets,” Sci. Rep., vol. 5, no. 1, pp. 1–9, Mar. 2015.

[11]	L. Guan, J. Tian, R. Cao, M. Li, Z. Cai, and W. Shen, “Barcode-like paper sensor for smartphone diagnostics: An application of blood typing,” Anal. Chem., vol. 86, no. 22, pp. 11362–11367, Nov. 2014.

[12]	M. Moazeni, F. Karimzadeh, and A. Kermanpur, “Peptide modified paper based impedimetric immunoassay with nanocomposite electrodes as a point-of-care testing of Alpha-fetoprotein in human serum,” Biosens. Bioelectron., vol. 117, no. June, pp. 748–757, 2018.

[13]	J. A. Adkins, K. Boehle, C. Friend, B. Chamberlain, B. Bisha, and C. S. Henry, “Colorimetric and Electrochemical Bacteria Detection Using Printed Paper- and Transparency-Based Analytic Devices,” Anal. Chem., vol. 89, no. 6, pp. 3613–3621, Mar. 2017.

[14]	D. M. Cate, J. A. Adkins, J. Mettakoonpitak, and C. S. Henry, “Recent developments in paper-based microfluidic devices,” Analytical Chemistry, vol. 87, no. 1. American Chemical Society, pp. 19–41, 06-Jan-2015.

[15]	E. W. Washburn, “The dynamics of capillary flow,” Phys. Rev., vol. 17, no. 3, pp. 273–283, Mar. 1921.

[16]	E. Fu, S. A. Ramsey, P. Kauffman, B. Lutz, and P. Yager, “Transport in two-dimensional paper networks,” Microfluid. Nanofluidics, vol. 10, no. 1, pp. 29–35, Jan. 2011.

[17]	C. K. Camplisson, K. M. Schilling, W. L. Pedrotti, H. A. Stone, and A. W. Martinez, “Two-ply channels for faster wicking in paper-based microfluidic devices,” Lab Chip, vol. 15, no. 23, pp. 4461–4466, Oct. 2015.

[18]	E. Elizalde, R. Urteaga, and C. L. A. Berli, “Rational design of capillary-driven flows for paper-based microfluidics,” Lab Chip, vol. 15, no. 10, pp. 2173–2180, May 2015.

[19]	S. Hong and W. Kim, “Dynamics of water imbibition through paper channels with wax boundaries,” Microfluid. Nanofluidics, vol. 19, no. 4, pp. 845–853, Oct. 2015.

[20]	M. P. Nguyen, N. A. Meredith, S. P. Kelly, and C. S. Henry, “Design considerations for reducing sample loss in microfluidic paper-based analytical devices,” Anal. Chim. Acta, vol. 1017, pp. 20–25, Aug. 2018.

[21]	N. Yang, C. Chen, P. Wang, J. Sun, and H. Mao, “Structure optimization method of microfluidic paper chip based on image grey-level statistics for chromogenic reaction,” Chem. Eng. Process. - Process Intensif., vol. 143, no. August, p. 107627, 2019.

[22]	W. Li et al., “Fully-drawn origami paper analytical device for electrochemical detection of glucose,” Sensors Actuators, B Chem., vol. 231, pp. 230–238, 2016.

[23]	J. Park and J. K. Park, “Pressed region integrated 3D paper-based microfluidic device that enables vertical flow multistep assays for the detection of C-reactive protein based on programmed reagent loading,” Sensors Actuators, B Chem., vol. 246, pp. 1049–1055, 2017.

[24]	J. Zhou et al., “ZnSe quantum dot based ion imprinting technology for fluorescence detecting cadmium and lead ions on a three-dimensional rotary paper-based microfluidic chip,” Sensors Actuators, B Chem., vol. 305, 2020.

[25]	W. J. Zhu et al., “Bienzyme colorimetric detection of glucose with self-calibration based on tree-shaped paper strip,” Sensors Actuators, B Chem., vol. 190, pp. 414–418, 2014.

[26]	W. T. Sow, F. Ye, C. Zhang, and H. Li, “Smart materials for point-of-care testing: From sample extraction to analyte sensing and readout signal generator,” Biosensors and Bioelectronics, vol. 170. Elsevier Ltd, p. 112682, 15-Dec-2020.

[27]	T. N.V.K.V. Prasad, V. Subba Rao Kambala, and R. Naidu, “A Critical Review on Biogenic Silver Nanoparticles and their Antimicrobial Activity,” Curr. Nanosci., vol. 7, no. 4, pp. 531–544, 2011.

[28]	M. Rycenga et al., “Controlling the synthesis and assembly of silver nanostructures for plasmonic applications,” Chem. Rev., vol. 111, no. 6, pp. 3669–3712, 2011.

[29]	W. Li, M. Chen, W. Li, C. You, J. Wei, and L. Zhi, “Synthesis of air stable silver nanoparticles and their application as conductive ink on paper based flexible electronics,” in Materials Research Innovations, 2014, vol. 18, pp. S4723–S4727.

[30]	X. C. Jiang and A. B. Yu, “Silver nanoplates: A highly sensitive material toward inorganic anions,” Langmuir, vol. 24, no. 8, pp. 4300–4309, Apr. 2008.

[31]	R. Ding, M. Fiedoruk-Pogrebniak, M. Pokrzywnicka, R. Koncki, J. Bobacka, and G. Lisak, “Solid reference electrode integrated with paper-based microfluidis for potentiometric ion sensing,” Sensors Actuators B. Chem., no. July, p. 128680, 2020.

[32]	D. C. M. Ferreira et al., “Optical paper-based sensor for ascorbic acid quantification using silver nanoparticles,” Talanta, vol. 141, pp. 188–194, Aug. 2015.

[33]	N. Ratnarathorn, O. Chailapakul, C. S. Henry, and W. Dungchai, “Simple silver nanoparticle colorimetric sensing for copper by paper-based devices,” Talanta, vol. 99, pp. 552–557, Sep. 2012.

[34]	A. Baranwal et al., “Chitosan: An undisputed bio-fabrication material for tissue engineering and bio-sensing applications,” International Journal of Biological Macromolecules, vol. 110. Elsevier B.V., pp. 110–123, 15-Apr-2018.

[35]	S. Parween et al., “Cross-linked chitosan biofunctionalized paper-based microfluidic device towards long term stabilization of blood typing antibodies,” Int. J. Biol. Macromol., vol. 163, pp. 1233–1239, Nov. 2020.

[36]	E. F. M. Gabriel, P. T. Garcia, T. M. G. Cardoso, F. M. Lopes, F. T. Martins, and W. K. T. Coltro, “Highly sensitive colorimetric detection of glucose and uric acid in biological fluids using chitosan-modified paper microfluidic devices,” Analyst, vol. 141, no. 15, pp. 4749–4756, 2016.

[37]	F. Li, X. Wang, J. Liu, Y. Hu, and J. He, “Double-layered microfluidic paper-based device with multiple colorimetric indicators for multiplexed detection of biomolecules,” Sensors Actuators, B Chem., vol. 288, no. August 2018, pp. 266–273, 2019.

[38]	L. Cao et al., “A disposable paper-based microfluidic immunosensor based on reduced graphene oxide-tetraethylene pentamine/Au nanocomposite decorated carbon screen-printed electrodes,” Sensors Actuators, B Chem., vol. 252, pp. 44–54, 2017.

[39]	C. Parolo, A. de la Escosura-Muñiz, and A. Merkoçi, “Enhanced lateral flow immunoassay using gold nanoparticles loaded with enzymes,” Biosens. Bioelectron., vol. 40, no. 1, pp. 412–416, Feb. 2013.

[40]	G. Palazzo, L. Facchini, and A. Mallardi, “Colorimetric detection of sugars based on gold nanoparticle formation,” Sensors Actuators, B Chem., vol. 161, no. 1, pp. 366–371, 2012.

[41]	F. Wang et al., “Paper-based closed Au-Bipolar electrode electrochemiluminescence sensing platform for the detection of miRNA-155,” Biosens. Bioelectron., vol. 150, no. November, 2020.

[42]	F. Figueredo, P. T. Garcia, E. Cortón, and W. K. T. Coltro, “Enhanced Analytical Performance of Paper Microfluidic Devices by Using Fe3O4 Nanoparticles, MWCNT, and Graphene Oxide,” ACS Appl. Mater. Interfaces, vol. 8, no. 1, pp. 11–15, 2016.

[43]	E. Evans, E. F. Moreira Gabriel, T. E. Benavidez, W. K. Tomazelli Coltro, and C. D. Garcia, “Modification of microfluidic paper-based devices with silica nanoparticles,” Analyst, vol. 139, no. 21, pp. 5560–5567, 2014.

[44]	X. Li, C. Zhao, and X. Liu, “A paper-based microfluidic biosensor integrating zinc oxide nanowires for electrochemical glucose detection,” Microsystems Nanoeng., vol. 1, pp. 1–7, 2015.

[45]	X. Sun et al., “Ultrasensitive microfluidic paper-based electrochemical/visual biosensor based on spherical-like cerium dioxide catalyst for miR-21 detection,” Biosens. Bioelectron., vol. 105, pp. 218–225, 2018.

[46]	J. Shen et al., “Bimetallic nano-mushrooms with DNA-mediated interior nanogaps for high-efficiency SERS signal amplification,” Springer.

[47]	Q. An, M. Yu, Y. Zhang, W. Ma, J. Guo, and C. Wang, “Fe 3O 4@carbon microsphere supported Ag-Au bimetallic nanocrystals with the enhanced catalytic activity and selectivity for the reduction of nitroaromatic compounds,” J. Phys. Chem. C, vol. 116, no. 42, pp. 22432–22440, Oct. 2012.

[48]	Z. Dehghani, M. Hosseini, J. Mohammadnejad, and M. R. Ganjali, “Novel colorimetric sensor based on peroxidase-like activity of chitosan-stabilized Au/Pt nanoclusters for trace lead,” Anal. Methods, vol. 11, no. 5, pp. 684–690, Feb. 2019.

[49]	M. Rahimi-Nasrabadi, F. Mizani, … M. H.-… A. P. A. M., and  undefined 2017, “Detection of hydrogen peroxide and glucose by using Tb2 (MoO4) 3 nanoplates as peroxidase mimics,” Elsevier.

[50]	X. Zhu, J. Huang, J. Liu, H. Zhang, J. Jiang, and R. Yu, “A dual enzyme-inorganic hybrid nanoflower incorporated microfluidic paper-based analytic device (μPAD) biosensor for sensitive visualized detection of glucose,” Nanoscale, vol. 9, no. 17, pp. 5658–5663, 2017.

[51]	A. Bagheri pebdeni and M. Hosseini, “Fast and selective whole cell detection of Staphylococcus aureus bacteria in food samples by paper based colorimetric nanobiosensor using peroxidase-like catalytic activity of DNA-Au/Pt bimetallic nanoclusters,” Microchem. J., vol. 159, p. 105475, Dec. 2020.

[52]	L. Wu et al., “A paper-based electrochemiluminescence electrode as an aptamer-based cytosensor using PtNi@carbon dots as nanolabels for detection of cancer cells and for in-situ screening of anticancer drugs,” Microchim. Acta, vol. 183, no. 6, pp. 1873–1880, Jun. 2016.

[53]	A. Arshad, H. Wang, X. Bai, R. Jiang, S. Xu, and L. Wang, “Colorimetric paper sensor for sensitive detection of explosive nitroaromatics based on Au@Ag nanoparticles,” Spectrochim. Acta - Part A Mol. Biomol. Spectrosc., vol. 206, pp. 16–22, Jan. 2019.

[54]	L. Deng, C. Chen, C. Zhu, S. Dong, and H. Lu, “Multiplexed bioactive paper based on GO@SiO2@CeO2 nanosheets for a low-cost diagnostics platform,” Biosens. Bioelectron., vol. 52, pp. 324–329, 2014.

[55]	Y. Fan et al., “A wireless point-of-care testing system for the detection of neuron-specific enolase with microfluidic paper-based analytical devices,” Biosens. Bioelectron., vol. 95, no. April, pp. 60–66, Sep. 2017.

[56]	B. Wei, K. Mao, N. Liu, M. Zhang, and Z. Yang, “Graphene nanocomposites modified electrochemical aptamer sensor for rapid and highly sensitive detection of prostate specific antigen,” Biosens. Bioelectron., vol. 121, no. August, pp. 41–46, 2018.

[57]	A. Manbohi and S. H. Ahmadi, “Sensitive and selective detection of dopamine using electrochemical microfluidic paper-based analytical nanosensor,” Sens. Bio-Sensing Res., vol. 23, no. February, p. 100270, 2019.

[58]	M. M. Liu et al., “MoOx quantum dots with peroxidase-like activity on microfluidic paper-based analytical device for rapid colorimetric detection of H2O2 released from PC12 cells,” Sensors Actuators, B Chem., vol. 305, no. November 2019, p. 127512, 2020.

[59]	F. H. Cincotto, E. L. Fava, F. C. Moraes, O. Fatibello-Filho, and R. C. Faria, “A new disposable microfluidic electrochemical paper-based device for the simultaneous determination of clinical biomarkers,” Talanta, vol. 195, pp. 62–68, 2019.

[60]	F. Auzel, “Upconversion and Anti-Stokes Processes with f and d Ions in Solids,” Chemical Reviews, vol. 104, no. 1. Chem Rev, pp. 139–173, Jan-2004.

[61]	Y. Tang, W. Gu, L. Fan, Y. Xu, F. Yin, and Y. Liu, “Test method study for SVG in wind farm and result analysis,” Gaoya Dianqi/High Volt. Appar., vol. 51, no. 6, pp. 15–19, Jun. 2015.

[62]	X. Chen et al., “A paper-supported aptasensor based on upconversion luminescence resonance energy transfer for the accessible determination of exosomes,” Biosens. Bioelectron., vol. 102, no. June 2017, pp. 582–588, 2018.

[63]	P. Jiang, M. He, L. Shen, A. Shi, and Z. Liu, “A paper-supported aptasensor for total IgE based on luminescence resonance energy transfer from upconversion nanoparticles to carbon nanoparticles,” Sensors Actuators, B Chem., vol. 239, pp. 319–324, 2017.

[64]	L. M. Fu and Y. N. Wang, “Detection methods and applications of microfluidic paper-based analytical devices,” TrAC - Trends Anal. Chem., vol. 107, pp. 196–211, 2018.

[65]	A. S. Tsagkaris, D. Migliorelli, L. Uttl, D. Filippini, J. Pulkrabova, and J. Hajslova, “A microfluidic paper-based analytical device (μPAD) with smartphone readout for chlorpyrifos-oxon screening in human serum,” Talanta, vol. 222, p. 121535, Jan. 2021.

[66]	X. Chen et al., “Determination of glucose and uric acid with bienzyme colorimetry on microfluidic paper-based analysis devices,” Biosens. Bioelectron., vol. 35, no. 1, pp. 363–368, 2012.

[67]	B. M. Jayawardane, I. D. McKelvie, and S. D. Kolev, “Development of a gas-diffusion microfluidic paper-based analytical device (μPAD) for the determination of ammonia in wastewater samples,” Anal. Chem., vol. 87, no. 9, pp. 4621–4626, 2015.

[68]	C. C. Tseng, R. J. Yang, W. J. Ju, and L. M. Fu, “Microfluidic paper-based platform for whole blood creatinine detection,” Chem. Eng. J., vol. 348, no. April, pp. 117–124, 2018.

[69]	T. Tian et al., “Integration of target responsive hydrogel with cascaded enzymatic reactions and microfluidic paper-based analytic devices (μPADs) for point-of-care testing (POCT),” Biosens. Bioelectron., vol. 77, pp. 537–542, Mar. 2016.

[70]	S. U. Son et al., “Naked-eye detection of pandemic influenza a (pH1N1) virus by polydiacetylene (PDA)-based paper sensor as a point-of-care diagnostic platform,” Sensors Actuators, B Chem., vol. 291, no. September 2018, pp. 257–265, 2019.

[71]	S. H. Im, K. R. Kim, Y. M. Park, J. H. Yoon, J. W. Hong, and H. C. Yoon, “An animal cell culture monitoring system using a smartphone-mountable paper-based analytical device,” Sensors Actuators, B Chem., vol. 229, pp. 166–173, 2016.

[72]	H. J. Chun, Y. M. Park, Y. D. Han, Y. H. Jang, and H. C. Yoon, “Paper-based glucose biosensing system utilizing a smartphone as a signal reader,” Biochip J., vol. 8, no. 3, pp. 218–226, 2014.

[73]	W. Y. Li et al., “Efficient in situ growth of enzyme-inorganic hybrids on paper strips for the visual detection of glucose,” Biosens. Bioelectron., vol. 99, pp. 603–611, 2018.

[74]	D. Calabria, C. Caliceti, M. Zangheri, M. Mirasoli, P. Simoni, and A. Roda, “Smartphone–based enzymatic biosensor for oral fluid L-lactate detection in one minute using confined multilayer paper reflectometry,” Biosens. Bioelectron., vol. 94, no. February, pp. 124–130, 2017.

[75]	K. Kaarj, P. Akarapipad, and J. Y. Yoon, “Simpler, Faster, and Sensitive Zika Virus Assay Using Smartphone Detection of Loop-mediated Isothermal Amplification on Paper Microfluidic Chips,” Sci. Rep., vol. 8, no. 1, p. 12438, Dec. 2018.

[76]	Monisha et al., “Inkjet-printed paper-based colorimetric sensor coupled with smartphone for determination of mercury (Hg2+),” J. Hazard. Mater., vol. 414, p. 125440, Jul. 2021.

[77]	M. de O. K. Franco, W. T. Suarez, V. B. dos Santos, I. S. Resque, M. H. dos Santos, and L. F. Capitán-Vallvey, “Microanalysis based on paper device functionalized with cuprizone to determine Cu2+ in sugar cane spirits using a smartphone,” Spectrochim. Acta - Part A Mol. Biomol. Spectrosc., vol. 253, p. 119580, May 2021.

[78]	“Electrochemical Methods: Fundamentals and Applications, 2nd Edition | Wiley.” [Online]. Available: https://www.wiley.com/en-us/Electrochemical+Methods%3A+Fundamentals+and+Applications%2C+2nd+Edition-p-9780471043720. [Accessed: 17-Aug-2020].

[79]	K. K. Hussain, J. M. Moon, D. S. Park, and Y. B. Shim, “Electrochemical Detection of Hemoglobin: A Review,” Electroanalysis, vol. 29, no. 10, pp. 2190–2199, 2017.

[80]	L. Cao, G. C. Han, H. Xiao, Z. Chen, and C. Fang, “A novel 3D paper-based microfluidic electrochemical glucose biosensor based on rGO-TEPA/PB sensitive film,” Anal. Chim. Acta, vol. 1096, pp. 34–43, 2020.

[81]	L. Cao, C. Fang, R. Zeng, X. Zhao, Y. Jiang, and Z. Chen, “Paper-based microfluidic devices for electrochemical immunofiltration analysis of human chorionic gonadotropin,” Biosens. Bioelectron., vol. 92, no. February, pp. 87–94, Jun. 2017.

[82]	A. Määttänen et al., “A low-cost paper-based inkjet-printed platform for electrochemical analyses,” Sensors Actuators, B Chem., vol. 177, pp. 153–162, 2013.

[83]	N. Ruecha, R. Rangkupan, N. Rodthongkum, and O. Chailapakul, “Novel paper-based cholesterol biosensor using graphene/polyvinylpyrrolidone/polyaniline nanocomposite,” Biosens. Bioelectron., vol. 52, pp. 13–19, 2014.

[84]	Y. Wang et al., “Label-free microfluidic paper-based electrochemical aptasensor for ultrasensitive and simultaneous multiplexed detection of cancer biomarkers,” Biosens. Bioelectron., vol. 136, no. February, pp. 84–90, 2019.

[85]	J. Noiphung, T. Songjaroen, W. Dungchai, C. S. Henry, O. Chailapakul, and W. Laiwattanapaisal, “Electrochemical detection of glucose from whole blood using paper-based microfluidic devices,” Anal. Chim. Acta, vol. 788, pp. 39–45, 2013.

[86]	G. Ruiz-Vega, T. Garcia-Berrocoso, J. Montaner, and E. Baldrich, “Paper microfluidics on screen-printed electrodes for simple electrochemical magneto-immunosensor performance,” Sensors Actuators, B Chem., vol. 298, no. May, p. 126897, 2019.

[87]	Y. Wang et al., “A novel label-free microfluidic paper-based immunosensor for highly sensitive electrochemical detection of carcinoembryonic antigen,” Biosens. Bioelectron., vol. 83, pp. 319–326, 2016.

[88]	S. Chaiyo et al., “Non-enzymatic electrochemical detection of glucose with a disposable paper-based sensor using a cobalt phthalocyanine–ionic liquid–graphene composite,” Biosens. Bioelectron., vol. 102, no. August 2017, pp. 113–120, 2018.

[89]	J. A. Adkins and C. S. Henry, “Electrochemical detection in paper-based analytical devices using microwire electrodes,” Anal. Chim. Acta, vol. 891, pp. 247–254, 2015.

[90]	G. Ruiz-Vega et al., “Electrochemical POC device for fast malaria quantitative diagnosis in whole blood by using magnetic beads, Poly-HRP and microfluidic paper electrodes,” Biosens. Bioelectron., vol. 150, p. 111925, 2020.

[91]	Z. Xu, Z. Liu, M. Xiao, L. Jiang, and C. Yi, “A smartphone-based quantitative point-of-care testing (POCT) system for simultaneous detection of multiple heavy metal ions,” Chem. Eng. J., vol. 394, p. 124966, Aug. 2020.

[92]	Y. Jiao et al., “3D vertical-flow paper-based device for simultaneous detection of multiple cancer biomarkers by fluorescent immunoassay,” Sensors Actuators, B Chem., vol. 306, p. 127239, 2020.

[93]	J. M. C. C. Guzman, L. L. Tayo, C. C. Liu, Y. N. Wang, and L. M. Fu, “Rapid microfluidic paper-based platform for low concentration formaldehyde detection,” Sensors Actuators, B Chem., vol. 255, pp. 3623–3629, 2018.

[94]	M. Ueland et al., “Capillary-driven microfluidic paper-based analytical devices for lab on a chip screening of explosive residues in soil,” J. Chromatogr. A, vol. 1436, pp. 28–33, 2016.

[95]	M. Liu et al., “An IoT-enabled paper sensor platform for real-time analysis of isothermal nucleic acid amplification tests,” Biosens. Bioelectron., vol. 169, p. 112651, Dec. 2020.

[96]	T. H. Ulep et al., “Smartphone based on-chip fluorescence imaging and capillary flow velocity measurement for detecting ROR1+ cancer cells from buffy coat blood samples on dual-layer paper microfluidic chip,” Biosens. Bioelectron., vol. 153, p. 112042, Apr. 2020.

[97]	A. K. Yetisen et al., “Paper-based microfluidic system for tear electrolyte analysis,” Lab Chip, vol. 17, no. 6, pp. 1137–1148, Mar. 2017.

[98]	A. Zamora-Gálvez, E. Morales-Narváez, J. Romero, and A. Merkoçi, “Photoluminescent lateral flow based on non-radiative energy transfer for protein detection in human serum,” Biosens. Bioelectron., vol. 100, pp. 208–213, 2018.

[99]	R. R. Anjana et al., “S,N-doped carbon dots as a fluorescent probe for bilirubin,” Microchim. Acta, vol. 185, no. 1, pp. 1–11, 2018.

[100]	P. G. Sutariya, H. Soni, S. A. Gandhi, and A. Pandya, “Single-step fluorescence recognition of As3+, Nd3+ and Br− using pyrene-linked calix[4]arene: application to real samples, computational modelling and paper-based device,” New J. Chem., vol. 43, no. 2, pp. 737–747, 2019.

[101]	X. Chen et al., “A novel naphthofluorescein-based probe for ultrasensitive point-of-care testing of zinc(II) ions and its bioimaging in living cells and zebrafishes,” Spectrochim. Acta - Part A Mol. Biomol. Spectrosc., vol. 229, p. 117949, Mar. 2020.

[102]	F. Li et al., “High-resolution temporally resolved chemiluminescence based on double-layered 3D microfluidic paper-based device for multiplexed analysis,” Biosens. Bioelectron., vol. 141, no. June, p. 111472, Sep. 2019.

[103]	W. Liu et al., “A molecularly imprinted polymer based a lab-on-paper chemiluminescence device for the detection of dichlorvos,” Spectrochim. Acta - Part A Mol. Biomol. Spectrosc., vol. 141, pp. 51–57, 2015.

[104]	W. Liu, J. Kou, H. Xing, and B. Li, “Paper-based chromatographic chemiluminescence chip for the detection of dichlorvos in vegetables,” Biosens. Bioelectron., vol. 52, pp. 76–81, 2014.

[105]	F. Li et al., “Multiplexed chemiluminescence determination of three acute myocardial infarction biomarkers based on microfluidic paper-based immunodevice dual amplified by multifunctionalized gold nanoparticles,” Talanta, vol. 207, p. 120346, Jan. 2020.

[106]	S. R. Chinnadayyala, J. Park, H. T. N. Le, M. Santhosh, A. N. Kadam, and S. Cho, “Recent advances in microfluidic paper-based electrochemiluminescence analytical devices for point-of-care testing applications,” Biosens. Bioelectron., vol. 126, pp. 68–81, Feb. 2019.

[107]	J. L. Delaney, C. F. Hogan, J. Tian, and W. Shen, “Electrogenerated chemiluminescence detection in paper-based microfluidic sensors,” Anal. Chem., vol. 83, no. 4, pp. 1300–1306, 2011.

[108]	Y. Chen, J. Wang, Z. Liu, X. Wang, X. Li, and G. Shan, “A simple and versatile paper-based electrochemiluminescence biosensing platform for hepatitis B virus surface antigen detection,” Biochem. Eng. J., vol. 129, pp. 1–6, 2018.

[109]	L. Chen, C. Zhang, and D. Xing, “Paper-based bipolar electrode-electrochemiluminescence (BPE-ECL) device with battery energy supply and smartphone read-out: A handheld ECL system for biochemical analysis at the point-of-care level,” Sensors Actuators, B Chem., vol. 237, pp. 308–317, 2016.

[110]	C. Liu, D. Wang, and C. Zhang, “A novel paperfluidic closed bipolar electrode-electrochemiluminescence sensing platform: Potential for multiplex detection at crossing-channel closed bipolar electrodes,” Sensors Actuators, B Chem., vol. 270, pp. 341–352, 2018.

[111]	K. Grudpan, S. D. Kolev, S. Lapanantnopakhun, I. D. McKelvie, and W. Wongwilai, “Applications of everyday IT and communications devices in modern analytical chemistry: A review,” Talanta, vol. 136. Elsevier B.V., pp. 84–94, 01-May-2015.

[112]	K. Yang, H. Peretz-Soroka, Y. Liu, and F. Lin, “Novel developments in mobile sensing based on the integration of microfluidic devices and smartphones,” Lab on a Chip, vol. 16, no. 6. Royal Society of Chemistry, pp. 943–958, 21-Mar-2016.

[113]	Z. Göröcs and A. Ozcan, “Biomedical imaging and sensing using flatbed scanners,” Lab Chip, vol. 14, no. 17, pp. 3248–3257, Sep. 2014.

[114]	X. Kou et al., “Smartphone-assisted robust enzymes@MOFs-based paper biosensor for point-of-care detection,” Biosens. Bioelectron., vol. 156, p. 112095, May 2020.

[115]	S. B. Puneeth and S. Goel, “Handheld and ‘Turnkey’ 3D Printed Paper-Microfluidic Viscometer with on-board Microcontroller for Smartphone based Biosensing Applications,” Anal. Chim. Acta, vol. 1153, p. 338303, Feb. 2021.

[116]	D. R. Bagal-Kestwal and B. H. Chiang, “Portable paper-micro well device composed of agglomerated nano-hematite clusters in enzyme-hydrogel composite for beta glucan detection using smartphone,” Sensors Actuators B Chem., p. 129836, Mar. 2021.

[117]	D. Kim, S. J. Kim, H. T. Ha, and S. Kim, “Smartphone-based image analysis coupled to paper-based colorimetric devices,” Curr. Appl. Phys., vol. 20, no. 9, pp. 1013–1018, Sep. 2020.

[118]	L. Huang et al., “Miniaturized Paper-Based Smartphone Biosensor for Differential Diagnosis of Wild-type Pseudorabies Virus Infection versus Vaccination Immunization,” Sensors Actuators, B Chem., vol. 327, p. 128893, Jan. 2021.

[119]	J. Xu and L. Yang, “A dual-readout paper-based sensor for on-site detection of penicillinase with a smartphone,” in MicroTAS 2020 - 24th International Conference on Miniaturized Systems for Chemistry and Life Sciences, 2020, vol. 335, pp. 522–523.

[120]	S. Arsawiset and S. Teepoo, “Ready-to-use, functionalized paper test strip used with a smartphone for the simultaneous on-site detection of free chlorine, hydrogen sulfide and formaldehyde in wastewater,” Anal. Chim. Acta, vol. 1118, pp. 63–72, Jun. 2020.

[121]	V. A. O. P. da Silva et al., “Microfluidic paper-based device integrated with smartphone for point-of-use colorimetric monitoring of water quality index,” Meas. J. Int. Meas. Confed., vol. 164, p. 108085, Nov. 2020.

[122]	Y. Feng et al., “A multifunctional fluorescent probe for visualizing H2S in wastewater with portable smartphone via fluorescent paper strip and sensing GSH in vivo,” J. Hazard. Mater., vol. 406, p. 124523, Mar. 2021.

[123]	H. Wang et al., “Colorimetric fluorescent paper strip with smartphone platform for quantitative detection of cadmium ions in real samples,” J. Hazard. Mater., vol. 392, p. 122506, Jun. 2020.

[124]	M. Maruthupandi, D. Thiruppathi, and N. Vasimalai, “One minute synthesis of green fluorescent copper nanocluster: The preparation of smartphone aided paper-based kit for on-site monitoring of nanomolar level mercury and sulfide ions in environmental samples,” J. Hazard. Mater., vol. 392, p. 122294, Jun. 2020.

[125]	T. H. Ulep and J. Y. Yoon, “Challenges in paper-based fluorogenic optical sensing with smartphones,” Nano Convergence, vol. 5, no. 1. Korea Nano Technology Research Society, p. 14, 01-Dec-2018.

[126]	J. H. Nichols, “Point-of-care testing,” in Contemporary Practice in Clinical Chemistry, Elsevier, 2020, pp. 323–336.

[127]	M. Rahman and A. B. Berenson, “Pregnancy Test Taking Is a Correlate of Unsafe Sex, Contraceptive Nonadherence, Pregnancy, and Sexually Transmitted Infections in Adolescent and Young Adult Women,” J. Women’s Heal., vol. 22, no. 4, pp. 339–343, Apr. 2013.

[128]	R. Nakhal, D. Wood, C. Woodhouse, and S. Creighton, “False-positive pregnancy tests following enterocystoplasty,” BJOG An Int. J. Obstet. Gynaecol., vol. 119, no. 3, pp. 366–368, Feb. 2012.

[129]	P. J. Watkinson, V. S. Barber, E. Amira, T. James, R. Taylor, and J. D. Young, “The effects of precision, haematocrit, pH and oxygen tension on point-of-care glucose measurement in critically ill patients: A prospective study,” Ann. Clin. Biochem., vol. 49, no. 2, pp. 144–151, Mar. 2012.

[130]	N. J. Perera, P. M. Stewart, P. F. Williams, E. L. Chua, D. K. Yue, and S. M. Twigg, “The danger of using inappropriate point-of-care glucose meters in patients on icodextrin dialysis,” Diabet. Med., vol. 28, no. 10, pp. 1272–1276, Oct. 2011.

[131]	C. Arnold, “At-home HIV test poses dilemmas and opportunities,” Lancet, vol. 380, no. 9847, pp. 1045–1046, Sep. 2012.

[132]	A. Manocha and S. Bhargava, “Emerging challenges in point-of-care testing,” Curr. Med. Res. Pract., vol. 9, no. 6, pp. 227–230, Nov. 2019.

[133]	M. Plebani, “Does POCT reduce the risk of error in laboratory testing?,” Clin. Chim. Acta, vol. 404, no. 1, pp. 59–64, Jun. 2009.

[134]	W. L. Then and G. Garnier, “Paper diagnostics in biomedicine,” Rev. Anal. Chem., vol. 32, no. 4, pp. 269–294, Nov. 2013.

[135]	A. Soni and S. K. Jha, “A paper strip based non-invasive glucose biosensor for salivary analysis,” Biosens. Bioelectron., vol. 67, pp. 763–768, 2015.

[136]	P. Basuri, A. Baidya, and T. Pradeep, “Sub-Parts-per-Trillion Level Detection of Analytes by Superhydrophobic Preconcentration Paper Spray Ionization Mass Spectrometry (SHPPSI MS),” Anal. Chem., vol. 91, no. 11, pp. 7118–7124, 2019.

[137]	J. F. da S. Petruci, P. C. Hauser, and A. A. Cardoso, “Colorimetric paper-based device for gaseous hydrogen cyanide quantification based on absorbance measurements,” Sensors Actuators, B Chem., vol. 268, pp. 392–397, 2018.

[138]	J. Xu et al., “Colorimetric and Electrochemiluminescence Dual-Mode Sensing of Lead Ion Based on Integrated Lab-on-Paper Device,” ACS Appl. Mater. Interfaces, vol. 10, no. 4, pp. 3431–3440, 2018.